PT 50

Part D 01 Calculation

Test Rig for Water Meters

Symbols and units	3

Measuring system - test rig	8

Virtual test system	10

Data for the testing process	11

1.	Temperatures for the test	11

2.	Pressure of the water in the test line	12

3.	Calculation of meter relative error	15

4.	Determination of conventional true value for weighing method of operation	18

5.	Flow	23

Formulas for calculations	25

Symbols and units

| Name | Description | Symbol | Symbol |
| --- | --- | --- | --- |
| Temp. UPAC | Actual temperature at the inlet | tupac | Tupac |
| Temp. DWAC | Actual temperature at the outlet | tdwac | Tdwac |
| Temp. DIAC | Actual temperature at the diverter | tdiac | Tdiac |
| Temp. UP ST | Temperature at the inlet at start of the test –first 8 measurements from the start | tupst | Tupst |
| Temp. UP EN | Temperature at the inlet at end of the test – last 8 measurement before end | tupen | Tupen |
| Temp. UP min | Minimum temperature at the inlet during measurement | tupmin | Tupmin |
| Temp. UP max | Maximum temperature at the inlet during measurement | tupmin | Tupmin |
| Temp. DW ST | Temperature at the outlet at start of the test - first 8 measurements from the start | tdwst | Tdwst |
| Temp. DW EN | Temperature at the outlet at end of the test - last 8 measurement before end | tdwen | Tdwen |
| Temp. DW min | Minimum temperature at the outlet during measurement | Tdwmin | Tdwmin |
| Temp. DW max | Maximum temperature at the outlet during measurement | Tdwmin | Tdwmin |
| Temp. DI ST | Temperature at the diverter / flowmeter at start of the test - first 8 measurements from the start | tdist | Tdist |
| Temp. DI EN | Temperature at the diverter / flowmeter at end of the test - last 8 measurement before end | tdien | Tdien |
| Temp. DI min | Minimum temperature at the diverter / flowmeter during measurement | Tdimin | Tdimin |
| Temp. DI max | Maximum temperature at the diverter / flowmeter during measurement | Tdimin | Tdimin |
| Temp. LN ST | Temperature in the test line at start of the test - first 8 measurements from the start - average from Temp. UP and Temp. DW | tdist | Tdist |
| Temp. LN EN | Temperature in the test line at end of the test - last 8 measurement before end - average from Temp. UP and Temp. DW | tdien | Tdien |
| Temp. LN min | Minimum temperature in the test line during measurement - average from Temp. UP and Temp. DW | Tdimin | Tdimin |
| Temp. LN max | Maximum temperature in the test line during measurement - average from Temp. UP and Temp. DW | Tdimin | Tdimin |
| Temp. stav | Average temperature at the start of the test | tstav | Tsta |
| Temp. enav | Average temperature at the end of the test | tenav | Tena |
| Temp. UP ME | Mean temperature at inlet | tupm | Tupm |
| Temp. DW ME | Mean temperature at outlet | tdwm | Tdwm |
| Temp. DI ME | Mean temperature at diverter | tdim | Tdim |
| Temp. LN ME | Mean temperature in the line | twame | Twame |
| Temp. AVG | Average temperature in the test line | tavg | Ta |
| Temp. dif-2 | Temperature difference - time | t | Tdt |

| Name | Description | Symbol | Symbol |
| --- | --- | --- | --- |
| Temp. dif-1 | Temperature difference – place | t | Tdp |
| Temp. Amb M | Ambient temperature mean | tambm | Tambm |
| Temp. Amb ST | Ambient temperature at the start of the measurement | tambst | Tambst |
| Temp. Amb EN | Ambient temperature at the end of the measurement | tamben | Tamben |
| Press. UP AC | Pressure at the inlet actual | pupac | Pupac |
| Press. DW AC | Pressure at the outlet actual | Pdwac | Pdwac |
| Press. UP ST | Pressure at the inlet at start of the test | pupst | Pupst |
| Press. UP EN | Pressure at the inlet at end of the test | pupen | Pupen |
| Press. UP min | Minimum pressure at the inlet of the test | pupmin | Pupmin |
| Press. UP max | Maximum pressure at the inlet of the test | pupmax | Pupmax |
| Press. DW min | Minimum pressure at the outlet of the test | pdwmin | Pdwmin |
| Press. DW max | Maximum pressure at the outlet of the test | pdwmax | Pdwmax |
| Press. DW ST | Pressure at the outlet at start of the test | pdwst | Pdwst |
| Press. DW EN | Pressure at the outlet at end of the test | pdwen | Pdwen |
| Press. UP avg | Average pressure at inlet | pupavg | Pupavg |
| Press. DW avg | Average pressure at outlet | pdwavg | Pdwavg |
| Press. UP ME | Mean pressure at inlet to the line | pupm | Pupm |
| Press. DW ME | Mean pressure at outlet from the line | pdwm | Pdwm |
| Press. AVG | Average pressure in the test line | pavg | Pa |
| Press. dif-2 | Pressure difference - time | p | Pdt |
| Press. dif 1 | Pressure difference – place | p | Pdp |
| Press. DE AC | Pressure difference measured by delta P meter – actual | PDEac | Pdeac |
| Press. DE ME | Pressure difference measured by delta P meter - mean | PDEME | Pdeme |
| Press. DE ST | Pressure difference measured by delta P meter – at the start of the test | PDEST | Pdest |
| Press. DE EN | Pressure difference measured by delta P meter – at the end of the test | PDEEN | Pdeen |
| Press. DE avg | Pressure difference measured by delta P meter – average | PDEavg | Pdest |
| Press. DE min | Pressure difference measured by delta P meter – minimum during the test | PDEmin | Pdeen |
| Press. DE max | Pressure difference measured by delta P meter – maximum during the test | PDEmax | Pdemax |
| Press. DE ME per mtr. | Pressure difference measured by delta P meter – mean during the test divided with number of tested meter (pressure difference per meter) | PDEMEm | Pdemem |
| Press. UP-DW ME | Pressure difference measured by pressure meter at the inlet and outlet - mean during the test | PDEUDME | Pdeudme |
| Press. UP-DW ME per meter | Pressure difference measured by pressure meter at the inlet and outlet - mean during the test divided with number of tested meter (pressure difference per meter) | PDEUDMEm | Pdeudmem |
| Name | Description | Symbol | Symbol |
| Press. PMax mean | Mean pressure for Pmax test | PmaxME | Pmaxme |
| Press. LN ME | Pressure in the line - mean | PLNME | Plnme |
| Press. LN ST | Pressure in the line start -at the start of the test | PLNST | Plnst |
| Press. LN EN | Pressure in the line - at the end of the test | PLNEN | Plnen |
| Press. LN avg | Pressure in the line average | PLNavg | Plnst |
| Press. LN min | Pressure in the line minimum value during the test | PLNmin | Plnen |
| Press. LN max | Pressure in the line maximum value during the test | PLNmax | Plnmax |
| Press. Amb M | Mean ambient pressure during tets | PambME | Pambme |
| Press. Amb ST | Ambient pressure at the start of the test | PambST | Pambst |
| Press. Amb EN | Ambient pressure at the end of the test | PambEN | Pamben |
| E rel. | Relative error of meter [ % ] |  | Er |
| E rel. Mt1 | Relative error of meter nr.1 in the line [ % ] |  | Er1 |
| E rel. Mt2 | Relative error of meter nr.2 in the line [ % ] |  | Er2 |
| E rel. Mti | Relative error of meter nr.i in the line [ % ] |  | Eri |
| Vol. Mt. | Volume measured by tested meter [ dm3 ] | VM | Vm |
| Vol. Mt.1 | Volume measured by tested meter nr.1 in the line [ dm3 ] | VM1 | Vm1 |
| Vol. rm | Reference volume for tested meter[ dm3 ] | VRM | Vrm |
| Vol. rm 1 | Reference volume for tested meter nr.1 in the line[ dm3 ] | VRM1 | Vrm1 |
| Vol. CTV | Reference volume conventional true value for the test [ dm3 ] | VR | Vr |
| Puls. Nrm | Number of pulses from master meter triggered from tested meter [ - ] |  | Nrm |
| Puls. Nrmac | Actual number of pulses from master meter triggered from tested meter during test [ - ] |  | Nrmac |
| Puls. Nr | Number of pulses from master meter triggered start and end of the test [ - ] |  | Nr |
| Puls. Nrac | Actual number of pulses from master meter triggered start, during test [ - ] |  | Nrac |
| Puls. Nm | Number of pulses from tested meter [ - ] |  | Nm |
| Puls. nr.i Nmi | Number of pulses from tested meter nr.i [ - ] |  | Nmi |
| Puls. Nmac | Actual number of pulses from tested meters, during test [ - ] |  | Nmac |
| Vol. RI | Volume obtained from master meter [ dm3 ] |  | Vri |
| Vol. W | Volume obtained from balance simple with balance correction and density correction [ dm3 ] |  | Vrb |
| Ma. R | Mass obtained from balance corrected [ kg] | MR | Mr |
| Km | Number of pulses per liter for tested meter [ -] | KM | Km |
| Km1 | Number of pulses per liter for tested meter nr.1. [ -] | KMi | Km1 |
| Kr | Number of pulses per liter for master meter [ -] | KR | Kr |
| Name | Description | Symbol | Symbol |
| Im | Number of liters per pulse for tested meter [ -] | IM | Im |
| Im1 | Number of liters per pulse for tested meter nr.1. [ -] | IMi | Im1 |
| Ir | Number of liters per pulse for master meter [ -] | IR | Ir |
| Irc | Number of liters per pulse for master meter corrected | IRC | Irc |
| Tau rt | Time of the test for conventional true volume [ s] | RT | Taurt |
| Tau m | Time of the test from first pulse to last pulse for the meter [ s] | m | Taum |
| Tau m1 | Time of the test from first pulse to last pulse for the meter nr.1 [ s] | m1 | Taum2 |
| Tau md | Test time from the tested meter obtained from data-stream from initial reading to end reading [ s] | md | Taumd |
| Tau md1 | Test time from the tested meter obtained from data-stream from initial reading to end reading for the meter nr.1 [ s] | md1 | Taumd1 |
| Tau mc | Test time from the tested meter obtained from camera from initial reading to end reading [ s] | mc | Taumc |
| Tau mc1 | Test time from the tested meter obtained from camera from initial reading to end reading for the meter nr.1 [ s] | mc1 | Taumc1 |
| T div (s) | Time of the test from gate start to end form diverter sensor [ s] | RB | Taurt |
| Diverter Start time | Time of the diverter switching from position 1 to position 2 in [ s], for flying start and stop | DI1 | Taudi1 |
| Diverter End time | Time of the diverter switching from position 2 to position 1 In [ s], for flying start and stop | DI2 | Taudi2 |
| Test time correction | Time correction obtained from diverter test doe determination systematic error | DI | Tauddi |
| Time between mass measurement | Time between Start measurement of weight and End measurement of weight - time for evaporation of water from the weighing tank | o | Tauo |
| Start valve open time | Opening time of the START-STOP valve in [ s] for standing start and stop | ZO | Tauzo |
| Start time close time | Closing time of the START-STOP valve in [ s] for standing start and stop | ZZ | Tauzz |
| Tau di | System delay mechanical - diverter | DI | Taudi |
| Ro AIRC | Density of the air during balance calibration [ kg/dm3 ] |  | Roairc |
| Ro WGC | Density of weights used for calibration [ kg/dm3 ] |  | Rowgc |
| Ro AIR | Density of air in the weighing tank during measurement [ kg/dm3 ] |  | Roair |
| Ro Wa | Water density flowing to the weighing tank corrected [ kg/dm3 ] |  | Rowa |

| Name | Description | Symbol | Symbol |
| --- | --- | --- | --- |
| Ro Wra | Water density in the weighing tank corrected [ kg/dm3 ] |  | Rowra |
| Kbuoyancy | Buoyancy value at measurement | KBU | Kbu |
| Mass ST | Mass corrected at the start of measurement [ kg ] |  | Mrst |
| Mass EN | Mass corrected at the end of measurement [ kg ] |  | Mren |
| Mass STB | Mass obtained from the weighing system at start [ kg ] |  | Mrstb |
| Mass ENB | Mass obtained from the weighing system at the end [ kg ] |  | Mrenb |
| Mass RB | Mass difference obtained from the weighing system [ kg ] |  | Mrb |
| Mass R | Mass difference obtained from the weighing system corrected [ kg ] |  | Mr |
| Ro Wat | Water density calculated from the temperature in the test line [ kg/dm3 ] |  | Rowat |
| Ro Cor | Density correction factor |  | Cro |
| Ro Watdim | Water density calculated from the density formula at temperature Tdim [ kg/dm3 ] – not corrected |  | Rowatdim |
| Ro Watupm | Water density calculated from the density formula at temperature Tupm [ kg/dm3 ] – not corrected |  | Rowatupm |
| Ro Watdwm | Water density calculated from the density formula at temperature Tdwm [ kg/dm3 ] – not corrected |  | Rowatdwm |
| Ro Warealt(r) | Measured density of sample water used at water density calibration at temperature (r) in [ kg/dm3 ] |  | Rowarealt(r) |
| Ro Wacalt(r) | Calculated density of water from table or formula at water density calibration temperature (r) [ kg/dm3 ] |  | Rowacalt(r) |
| Temp. Tcalt(r) | Temperature at which the density has been calibrated | t | Tro |
| Ro Wadim | Water density calculated from the density formula at temperature Tdim [ kg/dm3 ] – corrected |  | Rowadim |
| Ro Waupm | Water density calculated from the density formula at temperature Tupm [ kg/dm3 ] – corrected |  | Rowaupm |
| Ro Wadwm | Water density calculated from the density formula at temperature Tdwm [ kg/dm3 ] – corrected |  | Rowadwm |
| Kbuoyancy | Buoyancy correction during measurement | KBU | Kbu |
| Isother CF | Isothermal compressibility factor | B | B |
| Qm r | Flow calculated from conventional true mass and reference time [ 1000kg/h] |  | Qmb |
| Qrim | Flow obtained from master meter medium [ m3/h] |  | Qrim |
| Qctv | Flow obtained from Qrim with correction of master meter in [ m3/h] |  | Qctv |
| Qriac | Flow obtained from master meter actual [ m3/h] |  | Qriac |
| Flow i log | Logged flow from master meter | Qrilog | Qrilog |
| Name | Description | Symbol | Symbol |
| Cfq | Correction factor for flowmeter at flowrate |  | Cfq |
| Angle | Angle of the meter roller star or pointer from the start to the end of the measurement [deg] | m | Fim |
| Start angle | Starting angle at the measurement start [deg] | mST | Fimst |
| End angle | Angle at the measurement end [deg] | mEN | Fimen |
| Degree per liter | Pointer, star or roller value for volume calculation [deg/dm3] | Dl | Dpl |
| Qm r | Flow calculated from conventional true mass and time [ 1000kg/h] | Qmr | Qmr |
| Qref min | Minimum flow obtained from master meter during the test [ m3/h] | Qmrmin | Qmrmin |
| Qref max | Maximum flow obtained from master meter during the test [ m3/h] | Qmrmax | Qmrmax |
| Qref start | Flow at the start of the test obtained from master meter [ m3/h] - first 5 measurments | Qmrst | Qmrst |
| Qref end | Flow at the end of the test obtained from master meter [ m3/h] - last 5 measurements | Qmren | Qmren |
| Qctv min | Minimum flow obtained from master meter with correction of master meter in [ m3/h] | Qctvmin | Qctvmin |
| Qctv max | Maximum flow obtained from master meter with correction of master meter in [ m3/h] | Qctvmax | Qctvmax |
| Qctv start | Start flow obtained from master meter with correction of master meter in [ m3/h] - first 5 measurements | Qctvst | Qctvst |
| Qctv end | End flow obtained from master meter with correction of master meter in [ m3/h] - last 5 measurements | Qctven | Qctven |
| CQr | Correction factor for flowmeter | CQr | CQr |

Measuring system - test rig

The test rig is design for testing water meters and flowmeters in the wide flowrate range.

It is equipped with more adjusting lines, more feeding lines to reach the wide range with required specifications. Administrator of the Test Bench Framework (TBF) {program running on PC} can define the arrangement of the test rig according to required flowrate, temperature, pressure etc. This is enabling to use the optimum path for every in measurement. The test rig can have more test lines. In this case also the arrangement of test lines is definable. The test rig can work with 2 lines in serial connection or only using one line.

Adjusting lines are each for different flowrate range. More feeding lines can ensure stable flowrate in the wide flowrate range with good efficiency and conservation of energy.

With defined paths, TBF is calculating the conventional true value using sensors, weighing systems pressure meters, temperature meters using these components for calculation. TBF is collecting data from these components.

Test bench simplified functional scheme - screen from TBF:

We can explain the sensors in TBF with selected components as virtual test system.

For further data collection and calculation, TBF is considering only components in selected paths. Other components are ignore for certain measurement.

The following picture is demonstrating the active test path.

- Water Source. For this test P3 is active.VZ8 is opened.

- For Testing Table:   The 1st line is active. Active inlet valve is VL13 and active outlet valve is VL12. Active temperature sensors are T1, T2.  Active pressure sensors are P1 and P2

- Active valve for adjustment line is VI4. Active adjusting valve VR4.  Active diverter is DIV2 and active flowmeter is I3. Active weighing tank is WT2 and active outlet weighing tank valve is VB2. Active diverter temperature sensor is T6.

Temperature sensor T1 is at the inlet and is representing sensor Tup in virtual test system. Temperature sensor T2 is active at the outlet from the test line and is representing Tdw in virtual test system. Sensor T6 is active temperature sensor for measuring temperature at diverter DIV2 representing Tdi in virtual test bench.  Pressure sensor at the inlet to the test line Pr1 is active and is representing Pup and sensor Pr2 placed at the outlet from the test line is representing Pdw in the virtual test system. Reference point for calculation of conventional true value is situated in the middle between T1 T2 and between Pr1 and Pr2.

Virtual test system

In the virtual test system only active components are visible. In the following text the underline signs as “Tdi” are representing sensors.

These components are having certain role in the measuring process. Water temperature is measured at the inlet to the line with sensor Tup, then at the outlet from the line with sensor Tdw and at the DIVERTER with sensor Tdi. Pressure sensors used for this arrangement are Pup at the inlet to the test line and at the Pdw at the outlet from the test line.

Components of the virtual test bench

| Symbol | Description | Name in the text |
| --- | --- | --- |
|  | Valve at inlet STST valve | VLup |
|  | Valve at outlet | VLdw |
|  | Valve at adjustment branch outlet STST valve | Vinbr |
|  | Valve at adjustment branch outlet | Vend |
|  | Adjustment valve | Vreg |
|  | Temperature sensor at inlet | Tup |
|  | Temperature sensor at outlet | Tdw |
|  | Temperature sensor at diverter | Tdi |
|  | Temperature sensor at inlet to the test line | Pup |
|  | Pressure sensor at outlet from the test line | Pdw |
|  | Diverter | DIV |
|  | Diverter position sensor | Divs |
|  | Weighing measuring device | Balance |
|  | Outlet valve from weighing tank | Vb |
|  | Test line releasing valve at inlet | VLout1 |
|  | Test line releasing valve at outlet | VLout2 |
|  | Counter for tested meter | Nm |
|  | Counter for master meter trigger from meter | Nrm |
|  | Counter for master meter trigger from diverter | Nr |
|  | Start and stop valves used for standing start and stop. | STST |

# Data for the testing process

### Temperatures for the test

Temperatures are from following sensors of the virtual test bench

Temp. UP AC is from sensor Tup

Temp. DW AC is from sensor Tdw

Temp. DI AC is from sensor Tdi

Temp. ST 		 	[ oC]				(T1)

Temp. EN 		 	[ oC]				(T2)

Temp. dif-2		 		[ oC]				(T3)

Temp. dif-1				[ oC]				(T4)

Example of logged temperatures during test

| Timeactl | Tuplog | Tdwlog | Tdilog | l |
| --- | --- | --- | --- | --- |
| S | C | C | C | Number |
| 0.503 | 14.67 | 14.67 | 14.72 | 1 |
| 1,517 | 14.67 | 14.67 | 14.72 | 2 |
| 2,531 | 14.66 | 14.67 | 14.72 | 3 |
| 3,545 | 14.66 | 14.67 | 14.72 | 4 |
| 4,559 | 14.66 | 14.66 | 14.71 | 5 |
| 5,573 | 14.66 | 14.66 | 14.71 | 6 |
| 6,588 | 14.66 | 14.66 | 14.71 | 7 |
| 7,602 | 14.66 | 14.66 | 14.71 | 8 |
| 8,616 | 14.66 | 14.67 | 14.72 | 9 |
| 9,630 | 14.66 | 14.67 | 14.72 | 10 |

Temp. UP ME 			  	[ oC]					(T5)

Temp. DW ME 			 	[ oC]					(T6)

Temp. DI ME			 	[ oC]					(T7)

Temp. LN ME of water		  [ oC]					(T8)

Temp. LN ST				[ oC]					(T9)

Temp. LN EN				[ oC]					(T10)

For calculation of the temperature at the start of the test - first 8 measurements are taken to consideration. Test cycle is one s.

Temp. UP ST 			 [oC]					(T11)

Temp. DW ST 			[oC]						(T12)

Temp. DI ST 			 [oC]					(T13)

When standard PREMTHERM converter is in use for temperature measurement, the measuring cycle is 4 s. Practically it means  we will receive 4 same consequent values for each temperature.

For calculation of the temperature at the end of the test - last 8 measurements are taken to consideration. Test cycle is one s.

Temp. UP EN 			 [ oC]						(T14)

Temp. DW EN 			 [ oC]						(T15)

Temp. DI EN			 [ oC]						(T16)

### Pressure of the water in the test line

Values of pressures are from following sensors

Press. UP AC from sensor Pup

Press. DW AC from sensor Pdw

Press. UP avg			 [ kPa] 				(P1)

Press. DW avg 			[ kPa]				(P2)

Press. ST avg 			[ kPa]					(P3)

Press. EN avg 			[ kPa]					(P4)

Press. AVG			[ kPa]				(P5)

Press. Dif 2			[ kPa]					(P6)

Press. Dif 1			[ kPa]					(P7)

Press. DEEME 			[ kPa]					(P8)

For calculating pressure values TBF is using

Press. UP ME 			 	[kPa]					(P9)

Press. DW ME 			 	[kPa]					(P10)

Press. LN ME 				[kPa]					(P11)

Press. UP ST			 	[kPa]					(P12)

Press. DW ST				[kPa]					(P13)

Press UP EN				[kPa]					(P14)

Press DW EN				[kPa]					(P15)

Press. LN ME			 [kPa]					P16)

Press. LN ST 			 [kPa]					(P17)

Press. LN EN			 [kPa]					(P18)

Press. LN avg			 [kPa]					(P19)

### Calculation of meter relative error

Following values we receive from following sensors of the virtual test bench

Values for

- numer of pulses nRM are from counter nRM

- numer of pulses nR are from counter nR

- numer of pulses nM are from counter nM

#### Meter error

##### Relative error

E. Rel. 				  [ %]			(E1)

E rel. Mt1			 [ %]			(E2)

E rel. Mt1			 [ %]			(E3)

Similar for other meters

##### Error absolute

E. Abs.Mt			[ dm3]			(E4)

Abs. Mt1			[ dm3]			(E5)

Abs. Mt2			[ dm3]			(E6)

#### Volumes

##### Reference volume (correct volume) for synchronized method with master meter pulses

[ dm3]				(V1)

for Mt1 nr.1 			[ dm3]				(V2)

for Mt2 nr.2			[ dm3]				(V3)

Similar for other meters

##### Reference volume for synchronized method with time synchronization

TBF can also provide time synchronization method.

[dm3]				(V4)

for Mt1 nr.1 			  [dm3]				(V5)

for Mt2 nr.2			 [dm3]				(V6)

similar for other meters

##### Reference volume for synchronized method using datastram reading from tested meter

[dm3]				(V7)

for Mt1 nr.1			 [dm3]			(V8)

for Mt1 nr.2			 [dm3]				(V9)

##### Reference volume for synchronized method using process cameras

[dm3]			(V10)

for Mt1 nr.1			 [dm3]			(V11)

for Mt1 nr.2			 [dm3]			(V12)

##### Reference volume for methods without compensation and for standing start and stop

[dm3]				(V13)

for meter nr.1 			 [dm3]				(V14)

for meter nr.2			 [dm3]				(V15)

similar for other meters

##### Meter volume - Volume measured by meter

For method using pulses		  [ dm3] 		(V16)

for meter nr.1 				 [ dm3]		(V17)

for meter nr.2				 [ dm3] 	(V18)

##### Meter volume for data-stream method and for standing start and stop and method with process cameras

[ dm3]		(V19)

for meter nr.1 				[ dm3]		(V20)

for meter nr.2				 [dm3]		(V21)

##### Correct volume - conventional true volume

for weighing method 			 [dm3]			(V22)

Where

[dm3]		(V23)

for master meter method			[dm3]		(V24)

Where

Approximate volume from master meter	[dm3]		(V25)

Corrected volume from master meter		 [dm3]		(V26)

Corrected Irc value is given from		 [dm3]		(V27)

Where							     		(V28)

And  is function of Qrv  calculated from correction curve 		(V29)

### Determination of conventional true value for weighing method of operation

#### Volume of water flows to the weighing tank:

##### Conventional true volume

[ dm3] 				(V23)

[ -]					(M1)

In the current version of the TBF - buoyancy value is not calculated. It is a fix value written in TBF.

This value is set for each test bench between 1.00085 – 1.00106 depending on altitude and other conditions.

In near future we are planning to calculate the real buoyancy. All information for real buoyancy calculation are known.

System has information about ambient conditions. Density of weights we can obtain from calibration certificate of the weighing machines.

Because the temperature of water in the tank is unknown, for buoyancy calculation the water density   we can estimate as density at temperature tdim.

Mass RB   			 	[ kg]				(M2)

Mass ST			 [ kg]				(M3)

Mass EN			 [ kg]				(M4)

Where 		 correction of mass at   from calibration certificate

correction of mass at   from calibration certificate

Correction are obtain from balance calibration and interpolation between neighboring values.

Example.

e.g. At 15 kg the corrected value is 15.0035 kg.

#### Water density calculation

##### Density calculation formula: (Horst Betin, Frank Spieweck)

[kg/m3]			(R1)

| A0 | 9,99839520E+02 |
| --- | --- |
| A1 | 1,69525770E+01 |
| A2 | -7,99051270E-03 |
| A3 | -4,62417570E-05 |
| A4 | 1,05846010E-07 |
| A5 | -2,81030060E-10 |
| c | 1,68872360E-02 |

For this calculation, the temperature “t” is taken from temperature measurements.

[oC]							(R2)

[kg/m3]			(R3)

[kg/dm3]		(R4)

##### Pressure correction according to ISO 4064.

[kg/dm3]			(R5)

Where

##### Isothermal compressibility of water

[Pa-1]							(R6)

(R7)

| X0 | 5,08821e-10 |
| --- | --- |
| X1 | 1,2639418 |
| X2 | 0,2660269 |
| X3 | 0,3734838 |
| X4 | 2,0205242 |

##### Corrected water density to real water density

[kg/dm3]				(R8)

Where

[-]						(R9)

### Flow

Qm r 				[ m3/h]						(Q1)

Qm b 				[ kg/h]					(Q2)

Example of logged flowrates

| Timeactl | Qriac | l |
| --- | --- | --- |
| S | m3/h | Number |
| 0.503 | 97.54 | 1 |
| 1,517 | 97.56 | 2 |
| 2,531 | 97.61 | 3 |
| 3,545 | 97.74 | 4 |
| 4,559 | 97.72 | 5 |
| 5,573 | 97.55 | 6 |
| 6,588 | 97.59 | 7 |
| 7,602 | 97.81 | 8 |
| 8,616 | 97.83 | 9 |
| 9,630 | 97.56 | 10 |

Q rim for flying start and stop

[ m3/h]						(Q3)

Q rim for standing start and stop (without first 7 values or without first 30 % flow values from the test).

[ m3/h]						(Q4)

Qref start will be calculated as

[ m3/h] for flying start and stop			(Q5)

[ m3/h] for standing start and stop			(Q6)

Qref end will be calculated as

[ m3/h]					(Q7)

Flow given from master meter we calculate as

[ m3/h]					(Q8)

Correction factor for flowmeter is calculated as

[ m3/h]					(Q9)

Qctv start - will calculated as

[ m3/h]					(Q10)

Qctv end  - will be calculated as

[ m3/h]					(Q11)

# Formulas for calculations

#### The basic formula for calculation of the conventional true value for weighing method of operation is

[dm3]				(A1)

#### The basic formula for calculation of the conventional true value for master meter method of operation is

[dm3]							(A2)

#### Evaporation of water during the test for both methods

Evaporation during measurement from weighing tank

[kg]							(A3)

TBF is also storing o. This value is from PC clock not from PTB.

Mo - volume of water caused by evaporation of water during the test from the weighing tank

o  - is intensity of evaporation  in [l/s].

#### Diverter correction

Diverter correction is value which can be measured from the diverter test.

This is correction of the diverter for certain flowrate and for correct evaluation, the flowrate range of the diverter should be divide to 3 or more ranges. Each range should have different diverter correction.

Values between two flowrates are calculating using linear interpolation of two neighbouring values.

This corrections are explaining differences of the flow velocity profiles flowing from the diverter related to diverter position.

[-]						(A4)

#### Correction for calculating correct volume for tested meter using pulses from master meter

[-]							(A5)

#### Correction for calculating correct volume for tested meter using pulses time generator

[-]							(A6)

#### Correction for calculating correct volume for tested meter for datastrem method

#### [-]							(A7)

#### Correction for calculating correct volume for tested meter for method with process camera

[-]							(A8)

### Conventional true volumes

#### Conventional true volume for standing start and stop and weighing method of operation

Is the the same as basic conventional true volume described in 	(A1)

#### Conventional true volume for flying start and stop using weighing machines

[dm3] 	(A9)

#### Conventional true volume for flying start and stop using weighing machines prolonged test

#### [dm3]		(A10)

### Correct volumes for tested meter

#### Correct volume for tested meter for flying start and stop using weighing machines with pulse compensation from master meter and for prolonged method

[dm3]	(A11)

For standard method nRM is smaller as nR. For prolonged method nRM is bigger as nR.

#### Correct volume for tested meter for flying start and stop using weighing machines with pulse compensation from time generator

Secondary value

[dm3]	(A12)

#### Correct volume for tested meter for flying start and stop using weighing machines with data-stream system

[dm3]	(A13)

#### Correct volume for tested meter for flying start and stop prolonged using weighing machines with data-stream system

[dm3]	(A14)

#### Correct volume for tested meter for flying start and stop using weighing machines with process camera

[dm3]	(A15)

#### Correct volume for tested meter for flying start and stop prolonged using weighing machines with process camera

[dm3]		(A16)

#### Correct volume for tested meter for flying start and stop using master meter with pulse compensation from master meter

[dm3]					(A17)

#### Correct volume for tested meter for flying start and stop using master meter with pulse compensation time generating system

[dm3]					(A18)

#### Correct volume for tested meter for flying start and stop using master meter with datastream system

[dm3]					(A19)

#### Correct volume for tested meter for flying start and stop using master meter with process camera system

[dm3]					(A20)